# Calibration database: protocol for the two recording days

vMMR EEG rig — LSL marker timing correction (photodiode at the screen centre)

Prepared 13 October 2026. Computer B = PsychoPy stimulus PC; computer A = Simulink + g.HIamp recording PC. Background and results so far: docs/vMMR_marker_timing_calibration_2026-10-13.docx.

# 1. Purpose

Build a dataset large enough to estimate the marker-timing correction for the lab's ERP studies and to describe it in a methods section or supplement. Each run flashes a white square at the screen centre (where faces and primes appear) while an optical sensor on that square records the true onset with the EEG. The correction is the delay between each flash and its LSL marker placed by its LSL timestamp (pilot, 13 Oct: about 56 ms, SD 4.8 ms between runs).

The configuration is frozen for both days: nothing in the table below may change, and the sensor is not moved.

# 2. Frozen configuration

Table 1. Configuration for all database runs

| Item | Setting |
| --- | --- |
| Computer B script | run_photodiode_latency_database.py (from the repository, committed version) |
| Display (B) | Windows display mode 60 Hz (measures 59.94 Hz); nothing else changed |
| Square / sensor | 45 px white square at the screen centre; GTEC-0270 sensor taped over it for both days → g.TRIGbox → g.HIamp digital input |
| LSL | lsl_trigger.py, 600 samples/s (fixed in the script) |
| Simulink model (A) | matlab/on_eeg_computer/HIamp_record_LSL_DD_vMMR.slx with LSLMarkerTimestamps.m (R2020a, liblsl-Matlab on the path) |
| g.HIamp | 600 Hz, frame length 8 (the saved model was last used with 4: set 8 before day 1) |
| Simulink settings | fixed-step discrete 1/600 s, single-tasking, simulation pacing off, multithreaded co-simulation on (model on, g.HIamp block auto), workspace logging on, LSL Receive 600 Hz, stop constant 255 |
| Output folder (A) | E:\ACADEMIA\PhD\WP5\EEG_Recordings\Photodiode_Calib_vMMR\database |

# 3. Once, before day 1

Pull the committed repository on computer B (branch photodiode-latency-calibration); the script records the git commit in every run and git_tracked_changes: no confirms nothing was edited.

On computer A, open the model, set the g.HIamp frame length to 8, check the settings in Table 1 (Command Window: get_param(bdroot,'MultiThreadCoSim') = on; click the g.HIamp block, get_param(gcb,'MultiThreadCoSim') = auto) and save the model.

Create the output folder E:\ACADEMIA\PhD\WP5\EEG_Recordings\Photodiode_Calib_vMMR\database on computer A.

Fix the optical sensor at the screen centre with tape. Do one test run (not part of the database) and check that the g.TRIGbox LED follows the blinking square and that the end screen shows lsl_resyncs: 0.

Print the run log sheet (section 6) or keep data/database_run_log.csv open after each run.

# 4. Daily schedule (09:00–13:00, identical on both days)

Table 2. Day schedule

| Time | Activity | Runs |
| --- | --- | --- |
| 09:00–09:25 | Switch on / restart both computers (cold start). Check display 60 Hz, Simulink settings (Table 1, frame length 8), sensor position, output folder. | — |
| 09:25–10:55 | Block A: 12 short runs, one every ~7.5 min | 1–12 (run 1 = cold start) |
| 10:55–11:12 | Long run (500 flashes, ~12 min of flashes) | 13 |
| 11:12–11:25 | Restart both computers again (second cold start); copy block A files | — |
| 11:25–12:55 | Block B: 12 short runs | 14–25 (run 14 = cold start) |
| 12:55–13:00 | Copy all files to computer B and the backup drive; check the run log | — |

Per day: 24 short runs (100 flashes, ~2.6 min each plus ~3 min set-up) and 1 long run (500 flashes). If the day runs late, drop the last runs of a block rather than rushing the set-up. A run that is aborted or fails a check (section 7) is repeated as the next run number if time allows; it stays in the log. Minimum per day: 20 valid short runs and the long run.

# 5. Runs and file names

The script numbers runs per day automatically and shows the exact file names on the "start Simulink" screen; YYYYMMDD is that day's date. Only the run number (and short/long for run 13) changes between runs. Enter both names, with the folder, in the two Simulink To File blocks:

EEG: E:\ACADEMIA\PhD\WP5\EEG_Recordings\Photodiode_Calib_vMMR\database\photodiode_db_YYYYMMDD_<N>_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat

Timing: same name ending in _LSLtiming.mat

Table 3. Runs per day (same list on day 1 and day 2)

| Run | Start | Type | Cold start | EEG file (To File, main block) |
| --- | --- | --- | --- | --- |
| 1 | 09:25 | short | yes | photodiode_db_YYYYMMDD_1_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 2 | 09:32 | short |  | photodiode_db_YYYYMMDD_2_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 3 | 09:40 | short |  | photodiode_db_YYYYMMDD_3_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 4 | 09:47 | short |  | photodiode_db_YYYYMMDD_4_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 5 | 09:55 | short |  | photodiode_db_YYYYMMDD_5_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 6 | 10:02 | short |  | photodiode_db_YYYYMMDD_6_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 7 | 10:10 | short |  | photodiode_db_YYYYMMDD_7_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 8 | 10:17 | short |  | photodiode_db_YYYYMMDD_8_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 9 | 10:25 | short |  | photodiode_db_YYYYMMDD_9_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 10 | 10:32 | short |  | photodiode_db_YYYYMMDD_10_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 11 | 10:40 | short |  | photodiode_db_YYYYMMDD_11_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 12 | 10:47 | short |  | photodiode_db_YYYYMMDD_12_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 13 | 10:55 | long |  | photodiode_db_YYYYMMDD_13_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_long.mat |
| 14 | 11:25 | short | yes | photodiode_db_YYYYMMDD_14_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 15 | 11:32 | short |  | photodiode_db_YYYYMMDD_15_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 16 | 11:40 | short |  | photodiode_db_YYYYMMDD_16_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 17 | 11:47 | short |  | photodiode_db_YYYYMMDD_17_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 18 | 11:55 | short |  | photodiode_db_YYYYMMDD_18_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 19 | 12:02 | short |  | photodiode_db_YYYYMMDD_19_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 20 | 12:10 | short |  | photodiode_db_YYYYMMDD_20_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 21 | 12:17 | short |  | photodiode_db_YYYYMMDD_21_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 22 | 12:25 | short |  | photodiode_db_YYYYMMDD_22_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 23 | 12:32 | short |  | photodiode_db_YYYYMMDD_23_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 24 | 12:40 | short |  | photodiode_db_YYYYMMDD_24_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |
| 25 | 12:47 | short |  | photodiode_db_YYYYMMDD_25_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short.mat |

Timing file for each run: the same name with _LSLtiming before .mat, e.g. photodiode_db_YYYYMMDD_1_slx_frame_length_8_timestamped_LSL_trig_after_framerate_expstart_sqcenter_short_LSLtiming.mat.

Computer B writes, per run, into data/: db_YYYYMMDD_runNN_center_<type>_<time>_photodiode_latency_latency_flashes.csv, ..._run_info.txt, ..._frame_intervals.csv, ....log, and one row in data/database_run_log.csv.

# 6. Procedure for each run

B: start run_photodiode_latency_database.py. Dialog: operator, database_day (1 or 2), check the run number, run_type short (long for run 13), tick cold_start for runs 1 and 14, notes if anything unusual. OK.

B: the window opens and the refresh rate is measured (a few seconds). The screen then shows the two file names.

A: enter both file names in the To File blocks and start the Simulink model.

B: press SPACE. The alignment screen blinks the centre square: check the g.TRIGbox LED follows it, press SPACE.

B: experiment-start screen: press SPACE (sends marker 1). Calibration and flashes run by themselves (~2.6 min short, ~12 min long). Do not touch either computer.

B: end screen shows lsl_resyncs and dropped frames for 2 s; the script sends 255 and the Simulink model stops on its own.

A: check the two .mat files exist in the output folder. Write anything unusual in the log.

# 7. Checks and exclusion rules (fixed before data collection)

Table 4. A run is excluded if any rule applies

| Rule | How it is checked | When |
| --- | --- | --- |
| lsl_resyncs > 0 | end screen; run_info; run log | immediately |
| Run aborted or a file missing | run log completed; output folder on A | immediately |
| Dropped frames > 1 % of frames | end screen; run log | immediately |
| Backlog > 5 ms (stall the model never caught up) | notebook, LSL timestamp analysis | analysis |
| < 95 of 100 flashes paired (< 475 of 500) | notebook | analysis |

Excluded runs are counted and reported with the reason; they are not deleted. Runs failing an immediate check are repeated as the next run number.

# 8. End of each day

Copy E:\ACADEMIA\PhD\WP5\EEG_Recordings\Photodiode_Calib_vMMR\database (all .mat files) to computer B, data/EEG/database/.

Copy data/db_*, data/database_run_log.csv and the EEG files to the lab backup drive.

Note in the log: operator(s), room/lighting conditions, anything that differed from the protocol.

# 9. Analysis and expected precision

The compare notebook (notebooks/photodiode_latency_compare.ipynb) reads the database files (tags sqcenter, short/long). A database mode will apply Table 4 and report:

Table 5. What the database will report

| Quantity | Expected with ~48 short + 2 long runs (between-run SD ≈ 5 ms) |
| --- | --- |
| Correction (mean photodiode − timestamp-placed marker) | 95 % CI about ±1.3 ms |
| Between-session SD | estimated to about ±0.5 ms |
| Within-session SD (jitter) | pooled over ~5,800 flashes |
| 95 % prediction interval for one new session | about mean ± 9.5 ms |
| Drift within a session | 2 long runs: slope and first vs last 100 flashes |
| Effects of day, cold start, time of day | descriptive (2 days, 4 cold starts) |

Limitation to state in publications: two recording days sample day-to-day variation only twice. Re-check runs after any hardware or software change (3 runs) and periodic checks during data collection are added to the database as it grows.

# 10. After the database: to do

Copy all database files to computer B (data/EEG/database/, data/db_*, the run log) and back them up.

Calculate the correction: run the compare notebook in database mode (Table 4 exclusions); report the number of valid and excluded runs, the mean photodiode − timestamp-placed marker with its 95 % CI, the between- and within-session SD, the 95 % prediction interval for a new session, drift in the long runs, and day / cold-start effects. Freeze the correction (value, CI, date, git commit).

Implement the correction in the EEG pipeline: place each marker by its LSL timestamp (B timestamp + clock offset, mapped through A's clock to EEG samples) and add the frozen constant.

Update the experiment script run_vMMR_experiment_v1_trig_after_framerate_measurement.py: LSL keepalive and nominal rate 1200 → 600, experiment-start marker after the instructions, record the git commit and versions in run_info as the database script does.

Test with the experiment script: at least 3 full experiment sessions on different days with the database configuration, the timestamp model and the photodiode square at the centre on face-on frames. Report per face event the corrected marker against the photodiode onset (bias, SD, % within ±10 ms), lsl_resyncs = 0 and complete STD/DEV/TARGET codes. Then switch the photodiode square off for real sessions.

Ask Claude to write the final methods supplement for the article, starting from docs/methods_supplement_marker_timing_SKELETON.docx and filling in the database and validation results (all [TBD] fields).

Document the final configuration and correction in README.md and AGENTS.md; add re-check runs after any hardware or software change to the database.