# vMMR EEG rig: marker timing — update

What we did and decided on 8 and 13 October 2026, and where the code is

Repository: vMMR_EEG_PsychoPy_experiment. Computer B = PsychoPy stimulus PC (LSL sender); computer A = Simulink + g.HIamp recording PC (LSL receiver).

# Summary

The ~0.85 s delay between LSL markers and the real flash had one cause: PsychoPy pauses for ~0.8 s at the first flips of its fullscreen window. The LSL stream already existed then; the Simulink model takes one LSL sample per step, so it stalled with the stream and never caught up (the keepalive drops missed samples after pauses > 0.1 s). All EEG data then lagged the markers by ~0.8 s.

Fix: create the LSL stream after the window is open and the refresh rate is measured (*_v1_trig_after_framerate_measurement.py). Delay dropped from 864–891 ms to 33–63 ms; lsl_resyncs went from 1 to 0 in every run.

Correction: placing each marker by its LSL timestamp (saved by a new Simulink block) and adding ~56 ms gives the most stable result (SD 4.8 ms between runs); the recorded marker alone varies more (SD 7.8 ms) because its LSL transit varies between runs (1.6–12.3 ms).

Next: a two-day calibration database with the square at the screen centre, then the experiment script update and a final test.

# 8 October

Display: at the Windows 120 Hz setting PsychoPy only flipped at 60 Hz (AMD driver), so the display runs in 60 Hz mode; it measures 59.94 Hz, now the expected_refresh_hz default (frame counts are the same as at 60).

The count-based keepalive in lsl_trigger.py (commit "Fixed the LSL sync issue") was confirmed: 600.0 samples/s.

Frame-length runs (4, 6, 8) all showed 844–972 ms with no frame-length effect.

We added a second LSL inlet to the Simulink model, the MATLAB System block LSLMarkerTimestamps, which saves per model step: marker, rising-edge code, B timestamp, B→A clock offset, A's clock, samples pulled (*_LSLtiming.mat).

Four timestamp runs (baseline, PsychoPy started 1 min after Simulink, workspace logging off, multithreading off) showed: marker transit < 1 ms; one model stall of 0.80–0.84 s, 3.2 s before the START marker in every run; no LSL samples during it; lsl_resyncs: 1 on B. The PsychoPy log places it 0.06 s after getActualFrameRate starts. The Simulink changes made no difference.

Fix tested the same day: v1 (stream after the refresh measurement) and v2 (stream after the alignment screen) both removed the stall (delay 37.8 and 32.5 ms, lsl_resyncs: 0).

# 13 October

All latency scripts got an experiment-start screen after the sensor-alignment screen: SPACE sends marker 1 on the next flip, and the calibration and flashes are counted from it.

8 runs with v1, the timestamp model with its original settings, frame length 8 and 4 (4 each): no stall in any run; photodiode − recorded marker 39.8–62.9 ms; timestamp-placed 52.4–64.5 ms (mean 56.3, SD 4.8 between runs); within-run jitter ~5.5 ms SD; no difference between frame lengths 8 and 4.

The within-run jitter spans about one display frame and is identical for recorded and timestamp-placed markers, so it arises on the display side.

# Decisions

Table 1. Decisions

| Topic | Decision |
| --- | --- |
| LSL start-up | Create the LSL stream after the window and refresh measurement (v1); experiment-start marker 1 after the alignment screen |
| Simulink frame length | 8 |
| Simulink settings | original: multithreading on, workspace logging on, LSL Receive 600 Hz, timestamp block in every recording |
| LSL rate | 600 samples/s (keepalive and nominal) |
| Correction method | marker placed by its LSL timestamp + constant (from the database) |
| Within-run jitter (~5 ms SD) | accepted; no further display tuning |
| Database | 2 days × 09:00–13:00, square and sensor at the screen centre only, 24 short + 1 long run per day (≈ 50 sessions), protocol in docs/calibration_database_plan_2026-10.docx |

# Changes in the repository

Table 2. Files

| File | Change |
| --- | --- |
| run_photodiode_latency_v1_trig_after_framerate_measurement.py | LSL stream after the window/refresh measurement; on-screen Simulink prompt; experiment-start marker 1 |
| run_photodiode_latency_v2_trig_after_space_input.py | LSL stream after the alignment screen; marker 1 |
| run_photodiode_latency_v0.py | experiment-start screen and marker 1 added (13 Oct) |
| run_photodiode_latency_database.py | new: database runs (centre square, short/long, run log, versions and hashes in run_info) |
| run_vMMR_experiment_v1_trig_after_framerate_measurement.py | same LSL start-up fix for the experiment; 59.94 default; LSL statistics in run_info (keepalive default still 1200 → to change) |
| timing_utils.py, tests/ | 59.94 / 119.88 Hz accepted; tests |
| notebooks/photodiode_latency_compare.ipynb | new: multi-recording and LSL timestamp analysis |
| matlab/on_eeg_computer/ | the Simulink model with LSLMarkerTimestamps used on computer A |
| docs/ | calibration report (8 and 13 Oct versions), database protocol, jitter options, methods-supplement skeleton, this update |
| README.md | v1/v2, marker 1 and database sections |

Not in the repository: the recordings (data/ is git-ignored; they are on the lab computers) and lsl_api.cfg (network setting of computer B).

# Where the code is and how to get it

The work is on branch `photodiode-latency-calibration` of the fork https://github.com/DanielAtad/vMMR_EEG_PsychoPy_experiment (commit 3617721 from 8 October and the 13 October commit on top). It starts from the current main of omertzuk/vMMR_EEG_PsychoPy_experiment (0f8916e, "Fixed the LSL sync issue"), so it merges without conflicts as a fast-forward.

Option 1 — pull request on GitHub (recommended):

Open https://github.com/DanielAtad/vMMR_EEG_PsychoPy_experiment/pull/new/photodiode-latency-calibration

Set base repository omertzuk/vMMR_EEG_PsychoPy_experiment, base main; compare DanielAtad:photodiode-latency-calibration.

Review the files, then "Merge pull request".

Option 2 — command line:

git remote add daniel https://github.com/DanielAtad/vMMR_EEG_PsychoPy_experiment.git

git fetch daniel

# look at it first (optional)

git switch -c photodiode-latency-calibration daniel/photodiode-latency-calibration

# merge into main

git switch main

git pull

git merge --ff-only daniel/photodiode-latency-calibration

git push origin main

After merging, the lab computers can update with git pull on main. Computer A needs the files in matlab/on_eeg_computer/ (and liblsl-Matlab on the MATLAB path).

# Open points where your input would help

lsl_trigger.py: whether to send missed samples after a pause (raise _MAX_CATCHUP_S) so a future pause on computer B is recovered instead of shifting the rest of a run.

The experiment script: LSL rate 1200 → 600 and where its experiment-start marker goes (it now sends 9 after loading the stimuli; AGENTS.md lists 1 as experiment start).

The EEG pipeline function that places markers by their LSL timestamps and applies the correction.