# Within-run timing jitter: options if we decide to reduce it

vMMR EEG rig — prepared 13 October 2026 (decision that day: jitter accepted, not pursued)

Computer B = PsychoPy stimulus PC (AMD Radeon Vega 8 integrated graphics, Dell monitor, Windows 11); computer A = Simulink + g.HIamp recording PC.

# 1. What we know

Within a run, the delay between the marker and the photodiode onset varies with an SD of about 5.5 ms (all 13 October runs, frame length 8 and 4 alike).

The delays span about 19–20 ms (one 16.7 ms display frame plus 1.7 ms EEG sampling) and form two clusters about 11–12 ms apart.

It is the same for recorded markers and for markers placed by their LSL timestamps (r = 0.99) and unrelated to LSL transit (r ≈ 0), so it arises between the flip on computer B and the light reaching the sensor, not in LSL or Simulink.

Checked on the 13 October runs (800 flashes): unrelated to flash duration (r = −0.003), the preceding gap (r = −0.01), lock waiting in the LSL sender (r = 0), dropped frames (r = 0) and the phase of the push within the display frame (r = 0). Consecutive flashes fall in the same cluster only ~45 % of the time (chance level): the jitter is random from flash to flash.

Interpretation: the light onset is not locked to PsychoPy's flip. The most likely places are the Windows compositor / presentation path or the monitor's own processing (scaler, frame buffering, overdrive), both of which can add 0–1 frame of variable delay.

# 2. Is it worth reducing?

For ERP components the jitter acts like a small temporal smoothing. Treating a component as a Gaussian with width σc, jitter with SD σj widens it to √(σc² + σj²) and lowers its peak by the factor σc / √(σc² + σj²):

Table 1. Effect of jitter on peak amplitude (approximation)

| Component width σc | Peak amplitude with σj = 5.5 ms | With σj = 2 ms |
| --- | --- | --- |
| 10 ms (very sharp) | 88 % | 98 % |
| 20 ms (e.g. P1 / N170 flanks) | 96 % | 99.5 % |
| 40 ms (vMMN / P3-like) | 99 % | 99.9 % |

Mean latency is not biased by the jitter (it is removed by the correction); only the spread matters. For the planned vMMN analyses the loss is about 1–4 %, which is why it was accepted.

# 3. Important before changing anything

Any change to the display path changes the correction. If a change is adopted, the calibration database must be re-recorded with the new configuration (or the change must be made before the database).

Change one thing at a time, and keep everything else as in the database configuration.

Test each change with 3 short runs of run_photodiode_latency_database.py (centre square) plus the timestamp model; compare within-run SD, the cluster structure (histogram) and the mean delay with the 13 October values.

Acceptance criterion: within-run SD ≤ 2 ms with a single cluster, no lsl_resyncs, no increase in dropped frames.

# 4. Diagnostics (no hardware changes)

Identify the monitor model (the EDID code is DEL43D4) and its native refresh rate, response-time (overdrive) and low-latency / game modes from the on-screen menu.

Check whether the monitor is driven at its native rate: at the Windows 120 Hz setting PsychoPy could only flip at 60 Hz, which suggests a frame-rate conversion path somewhere.

Run one test with a high-speed camera or a second photodiode at the top of the screen to see whether the two clusters are present at all screen positions.

# 5. Changes to test, in order of effort

Table 2. Candidate changes

| # | Change | Where | Why |
| --- | --- | --- | --- |
| 1 | Disable "fullscreen optimizations" for the PsychoPy python.exe; Windows Game Mode off | Windows: python.exe Properties → Compatibility; Settings → Gaming | Forces the classic exclusive fullscreen path instead of the compositor |
| 2 | Hardware-accelerated GPU scheduling off (then on), reboot | Windows Settings → Display → Graphics | Changes the presentation queue |
| 3 | AMD Software: FreeSync off, Enhanced Sync off, Wait for Vertical Refresh "Always on", Radeon Anti-Lag / Chill / Frame Rate Target Control off | AMD Software (Adrenalin) → Gaming / Display | Removes variable-refresh and frame-pacing features |
| 4 | Monitor: overdrive/response time "Normal", dynamic contrast off, low-latency / game mode on | Monitor on-screen menu | Monitor processing can add 0–1 frame of variable delay |
| 5 | PsychoPy: core.rush(True) during flashes; try winType='glfw'; check waitBlanking | script | Reduces scheduling effects; different presentation back-end |
| 6 | Run the display at 120 Hz with PsychoPy flipping every frame (expected_refresh_hz 119.88) | Windows + AMD settings; needs the 60 Hz-only flip issue solved | Halves the frame quantum: jitter SD would drop to ~2.7 ms even if the asynchrony remains |
| 7 | Dedicated GPU instead of the integrated Vega 8, or a monitor with documented low, fixed input lag | hardware | Most likely to give stable, single-cluster timing |

# 6. Estimated effort

Steps 1–5: about half a day (3 runs ≈ 20 min per change, plus reboots). Step 6: half a day to a day, depending on the driver. Step 7: purchase and a full re-validation. Any adopted change requires the calibration database to be recorded again with the new configuration.