# vMMR EEG: marker timing – how to record and how to correct

Lab guide, 15 October 2026. Repository: https://github.com/DanielAtad/vMMR_EEG_PsychoPy_experiment (branch photodiode-latency-calibration). Methods and results: Supplementary Material S1 (docs/supplement_marker_timing_calibration_2026-10-15.docx).

# 1. In one paragraph

Event markers are sent by the PsychoPy computer (B) over LSL and recorded with the EEG on the g.HIamp/Simulink computer (A). The marker channel inside the EEG file is not accurate enough for ERP timing (it varies by about 9 ms between sessions). Instead, every event is placed by its LSL send timestamp, which a second LSL block saves in a separate timing file, and a calibrated constant is added: onset = timestamp-placed marker + 62.4 ms. This was measured with an optical sensor in 22 sessions (between-session SD 1.4 ms; single events SD 5.7 ms, all within one screen frame). The constant is valid only for the configuration below and for stimuli at the screen centre.

# 2. Computer B – stimulus PC (PsychoPy)

| Item | Required setting |
| --- | --- |
| Display | Dell P2426H, 1920×1080. Windows: Settings → System → Display → Advanced display → refresh rate 59.94 Hz (not 60, not 120). Check before each session day: .venv\Scripts\python check_display_refresh_rate.py must report ≈ 59.94 Hz and 100 % of intervals in 14–19 ms. |
| Software | Project .venv: Python 3.10.22, PsychoPy 2026.2.4, pylsl 1.18.5. Do not update without re-calibrating. |
| Network | lsl_api.cfg in the project folder with KnownPeers = {192.168.50.1} (computer A). |
| Experiment script | run_vMMR_experiment_v1_trig_after_framerate_measurement.py (until the updated version from the to-do document exists). Dialog: expected_refresh_hz 59.94, lsl_keepalive_hz 600, lsl_nominal_srate 600, send_LSL_triggers on, photodiode_square off for real participants (the current default is on). |
| Files needed | run_vMMR_experiment_v1_trig_after_framerate_measurement.py, lsl_trigger.py, timing_utils.py, conditions/, stimuli/, check_display_refresh_rate.py, lsl_api.cfg. |
| After every session | Open the newest data/*_run_info.txt: lsl_resyncs: 0 is required. Keep all outputs (_events.csv, _trials.csv, _run_info.txt, _frame_intervals.csv, .log). |

# 3. Computer A – recording PC (g.HIamp, Simulink)

| Item | Required setting |
| --- | --- |
| Software | MATLAB/Simulink R2020a with g.HIsys; liblsl-Matlab on the MATLAB path (lsl_loadlib must work). |
| Model | matlab/on_eeg_computer/HIamp_record_LSL_DD_vMMR.slx with LSLMarkerTimestamps.m in the same folder (or on the path). Do not use the older .slx files in the repository root. |
| g.HIamp block | 600 Hz, frame length 8 (block configuration dialog). |
| Solver / model settings | Fixed-step, discrete, step 1/600, single-tasking (“treat each discrete rate as a separate task” unchecked), simulation pacing off, multithreaded co-simulation on (model on, g.HIamp block auto), workspace logging on. |
| LSL blocks | LSL MARKERS FROM EXPERIMENT (g.HIsys LSL Receive): stream experiment_markers, sampleRate 600. LSL MARKER TIMESTAMPS (MATLAB System, interpreted execution): StreamName experiment_markers, ResolveTimeout 5, CorrectionInterval 1. |
| Output files | Two To File blocks, same base name: <name>.mat (variable EEG_data) and <name>_LSLtiming.mat (variable LSL_timing, Timeseries). Rename both before every session; the timing file is required for the correction. |
| Stop | The model stops on marker 255. The current experiment script ends with 99, so stop the model manually after the end screen until the script is updated. |

# 4. Session order

B: start the experiment script, fill in the dialog. The window opens and the refresh rate is measured.

B: the LSL stream is created and the screen asks to start the recording model.

A: set both output file names, then start the Simulink model (the stream must already exist: the timestamp block looks for it at model start, for 5 s).

B: the script continues when Simulink is connected; run the experiment.

End: stop the model (see above), check lsl_resyncs: 0 on B, copy the two .mat files from A and all outputs from B to the study folder.

# 5. Applying the correction (analysis)

analysis/marker_timing.py does the whole correction for one recording. It needs numpy, pandas and h5py (the .mat files are MATLAB v7.3/HDF5).

import sys; sys.path.insert(0, "analysis")    # repository analysis folder

from marker_timing import correct_events, to_mne_events

# the ..._LSLtiming.mat file must be next to the EEG file

events, qc = correct_events("main_..._sub01.mat",

run_info="sub01_..._run_info.txt")

if not qc["passed"]:

print("CHECK THIS SESSION:", qc["problems"])

events[["code", "onset_s", "sample"]]         # corrected time (s) and EEG row

mne_events = to_mne_events(events, codes=[11, 12, 21, 22, 31, 32, 41, 42])

Or from the command line (writes <recording>_events_corrected.csv):

.venv\Scripts\python analysis\marker_timing.py main_..._sub01.mat ^

--run-info sub01_..._run_info.txt

| Column / field | Meaning |
| --- | --- |
| code | marker code (11–44 face events, 9 start, 99 end, …) |
| onset_s | corrected onset on the EEG time axis (column 0 of EEG_data), in seconds |
| sample | row of EEG_data closest to the corrected onset (use as the event sample) |
| recorded_s, shift_vs_recorded_ms | where the marker channel had the event, and the shift applied |
| ts_B | LSL send timestamp on computer B (matches the PsychoPy LSL timestamp of the event) |
| qc | passed, problems, lsl_resyncs, lasting_backlog_ms, stalls_over_100ms, clock drift |

Matching events to the PsychoPy events file: face events (codes 11–44) appear in the same order as the rows of *_events.csv; check that the code sequences are identical. After the planned script update the events file will also contain each event's LSL timestamp (ts_B), which allows exact matching.

Practice trials send the same codes; drop them using the events file (phase).

Session check: use only sessions with qc["passed"] (lsl_resyncs 0, no lasting backlog of the model).

Stimulus position: 62.4 ms is for the screen centre (faces). The prime word (260 px above the centre) appears about 4 ms earlier because the screen is drawn top to bottom (estimate, not measured).

Never add 62.4 ms to the marker channel of EEG_data; the constant belongs to the timestamp method only.

# 6. When to re-calibrate

After any change of monitor, graphics driver, Windows display settings, PsychoPy/Python packages, lsl_trigger.py, the Simulink model or g.HIamp settings: record at least 3 calibration sessions with run_photodiode_latency_v1_trig_after_framerate_measurement.py (optical sensor at the screen centre) and analyse them with analysis/calibration_database_analysis.py. If they fall outside 59.4–65.3 ms, a new constant is needed.

# 7. Where things are

| What | Where |
| --- | --- |
| Correction function | analysis/marker_timing.py |
| Calibration analysis | analysis/calibration_database_analysis.py |
| Calibration data | Google Drive: final database/ (EEG_Computer_A/, Computer_B/, run_manifest.csv) |
| Recording model (A) | matlab/on_eeg_computer/ |
| Methods supplement | docs/supplement_marker_timing_calibration_2026-10-15.docx / .pdf |
| Open work | docs/todo_experiment_and_analysis_2026-10.docx |